Plant Petiole Nitrate Test Kit Instructions

Introduction

This kit contains everything you need to test plant petiole samples for nitrate content.  We have provided in this kit a 500 ppm Nitrate-N standard (D).  This is to ensure that your kit is working properly and to use as a reference tool when analyzing the color of your samples.

After comparing your samples with the nitrate standard and the color chart, you will have a value that determines the nitrate content of your plant sample.  Although we cannot guarantee the precision that you would expect from a testing lab, you will get accurate results immediately about the nitrate content of your plants that you can use for making preliminary decisions about nutrient application.

Kit Contents (per 5 samples):

- 1 cardboard cuvette holder

- 5 clear collection tubes with white screw caps containing 5 mL water

- 5 clear twist off squeeze bulbs (contains buffer)

- 1 clear snip off squeeze bulb containing nitrate standard

- 6  square reaction cuvettes in sealed foil pouch (contains enzyme)

- 10  plastic pipettes (for transferring one drop of liquid to containers)

- 6  clear glass tubes of color reagent powder (in bubble wrap bag)

- 5 plastic pans for plant sap collection

Materials you will need:

- Marker for labeling tubes and cuvette caps

- Scissors (to remove tip from snip off squeeze bulb)

- Pen (for recording data on data sheet)

- Garlic press, clean pliers, or similar apparatus to extract plant sap

DO:

- Store this kit in a cool and dark place (below 72°F/22°C), refrigerate if possible

- Follow the directions of a Certified Crop Advisor for advice on plant sampling.  A common method for plant testing calls for extracting sap from the petioles.  Contact NECi for more information.

- Run tests and standard all at once to ensure accuracy

- Label collected sample tubes and reaction cuvettes accordingly

- Gently shake cuvettes several times throughout reaction without inverting them.

DO NOT:

- Open sealed packets or mix cuvette contents until ready to use

- Add more than one drop of sample water to reaction cuvette

- Compare samples to color chart before waiting at least 10 minutes

- Invert reaction cuvettes when mixing or get liquid in the cap portion before step 4

Procedure

Step 1						Sample Preparation                                                                                                                          `

- Use a garlic press or clean pair of pliers to squeeze several drops of plant juice into plastic pan. (1.1)

- Repeat for each plant sample, using a clean pan for each.  Clean press or pliers between samples.

- Add one drop of plant juice to each tube with water using a new pipette for each sample. (1.2)  Label Tubes

- Recap the tubes and shake each tube for approximately 1 minute.

*Be sure to label each cap and record the information regarding the samples on the data sheet

Step 2						Preparing the Reaction Cuvettes                                                                                                    `

- Remove the 6 square reaction cuvettes from the foil pouch and place them in the cardboard cuvette holder,                          keeping the one with the blue dot separate from the others (this is your nitrate standard).

- Tap cuvettes to settle contents making sure to keep them upright to not lose contents.

- Snip off the end of the standard squeeze bulb and empty the entire contents into the cuvette with the blue dot.     This is the nitrate standard. (2.3)

- Label the cuvette caps with numbers, letters, or location. (1-5, A-E, etc.)

- Into each of the other 5 cuvettes, empty entire contents of one twist off squeeze bulb. (2.5)

- Cap all cuvettes and mix gently.  Do not invert (keep liquid out of cap).  Step 3 must follow within 5 minutes.

Step 3		 					Adding samples to Reaction Cuvettes                                                                                        `

- Within 5 minutes, pick up liquid from near the surface of your settled soil sample with a plastic pipette (3.1).

- Transfer only one drop (approximately 50 µl) of this liquid to the corresponding cuvette (Sample 1 into cuvette 1, etc.) making sure not to touch the sides or surface of the liquid with the pipette.  (3.2)

- Repeat for the remaining 4 samples, making sure to use a new pipette for each unique sample.

- Recap the cuvettes and mix by gently shaking side to side.

- Let set for at least 10 minutes, mixing gently every few minutes.

Step 4		 					Color development                                                                                                                                    `

- Tap each glass tube to settle powder; snap off top at score line in neck (4.1).

- Add color reagent powder to all reaction cuvettes, including the nitrate standard. (4.2)

- Firmly seal the cuvettes with square caps and vigorously mix each cuvette rapidly.

- Let cuvettes develop color for about 5 minutes, mixing them several times to dissolve most of the powder. Some powder may settle to the bottom of tube; this is okay.

Step 5							Evaluating your results                                                                                                                  `

After color development, compare your results within one hour using the nitrate standard and this chart or download NECi’s free color slider application for mobile devices.  The nitrate standard is set at 500 parts per million Nitrate-N (2200 ppm Nitrate).   Contact your local crop advisor for nutrient application advice.

Nitrate-N (ppm)                 0                         200                        500                          1000

Nitrate (ppm)                     0	                  900                       2200                         4400

Visit http://nitrate.com/petiolesapnitratetesting for more information about petiole testing

# Text boxes

D

C

B

A

H

G

F

E

Search for “Nitrate Color Slider” in your application store to download NECi’s free app.  Tap the screen once downloaded to toggle to the “plant” screen.  Slide the white circle along the line below the color block to analyze your sample.

NECi’s Color Slider is also available on your web based browser at www.nitrate.com/nitratecolorslider